Notas GPIB
Pin #	Signal names	Signal Description	Pin #	Signal names	Signal Description
1	DIO1	Data Input/Output Bit 1	13	DIO5	Data Input/Output Bit 5
2	DIO2	Data Input/Output Bit 2	14	DIO6	Data Input/Output Bit 6
3	DIO3	Data Input/Output Bit 3	15	DIO7	Data Input/Output Bit 7
4	DIO4	Data Input/Output Bit 4	16	DIO8	Data Input/Output Bit 8
5	EIO	End-Or-Identify	17	REN	Remote Enable
6	DAV	Data Valid	18	Shield	Ground (DAV)
7	NRFD	Not Ready For Data	19	Shield	Ground (NRFD)
8	NDAC	Not Data Accepted	20	Shield	Ground (NDAC)
9	IFC	Interface Clear	21	Shield	Ground (IFC)
10	SRQ	Service Request	22	Shield	Ground (SRQ)
11	ATN	Attention	23	Shield	Ground (ATN)
12	Shield	Chassis Ground	24	Single GND	Single Ground

GPIB On-Line Standards and Specifications Information

ANSI/IEEE 488.1 [IEC 60625-1] IEEE Standard Digital Interface for Programmable Instrumentation
Defines the Physical and Electrical layer, and signaling protocol.

ANSI/IEEE 488.2 IEEE Codes, Formats, Protocols, and Common Commands, and Standard Commands for Programmable Instruments

IEC/IEEE 60488-2 defines communication protocols for devices connected via IEEE 488 buses and common commands and characteristics useful in a wide range of instrument applications. It includes message-handling protocols, status reporting structures, and system configuration and synchronization protocols.

HS488 High-Speed GPIB Handshake Protocol Hand shake protocol, increases bus transfers to 8MBytes/s with other HS488 devices. NDAC is not required to Handshake, The Talker outputs data with DAV, waits then outputs new data with DAV without NDAC occurring.

SCPI Standard Commands for Programmable Instruments Defines a programming language for Instruments

IEEE488 Bus Interface IC Manufacturers

Standard TTL levels (2.0v/0.8v) @ 5.2mA source / 48mA sink, Totem Pole, Open Collector and Tristate devices are used with IEEE488.
Open Collector: SRQ, NRFD, NDAC
Open Collector or Tristate: ATN, IFC, REN, EOI, DAV.

MC3446; Quadruple Bus Transceiver [IEEE Standard 488 Compliant], 16-pin DIP [Obsolete part].
MC3447; Bidirectional Instrumentation Bus [GPIB] Transceiver, 24-pin DIP [Obsolete part].
TI {Controller TMS9914A-Transceiver, 75xx160/161/162}
Texas Instruments: DS75160A,DS75161A

There is not much special about driving over a IEEE488 bus, the original interface ICs were just normal TTL devices.
The Octal Transceiver 74F588 was used in many applications.

IEEE-STD-488 I/O Characteristics
Single Type	Digital Value
Input Voltage High:	VIH = 3.4 volts typical, 2.4 volts minimum
Input Voltage Low:	VIL = 0.22 volts typical, 0.4 volts maximum
Input Current High:	IIH = 2.5mA maximum
Input Current Low:	VIL = -3.2mA maximum
Output Voltage High:	VOH = 3.4 volts typical, 2.5 volts minimum
Output Voltage Low:	VOL = 0.22 volts typical, 0.5 volts maximum
Output Current High:	IOH = -5.2mA maximum
Output Current Low:	IOL = 48mA maximum

IEEE488 Data Bus Transfer Timing

The IEEE488 bus operates at the speed of the slowest device, all devices have to be ready before operation begins.
Signals are active low. Data is transferred asynchronous, using the Handshake lines instead of a clock.
Only Parallel Polling does not use the Handshake, all other transfers use Hand-shaking.

HS488 High-Speed GPIB Handshake Protocol Hand shake protocol [not shown above], increases bus transfers to 8MBytes/s with other HS488 devices.
NDAC is not required to Handshake, the Talker outputs data with DAV, waits then outputs new data with DAV without NDAC occurring.

GPIB Cable Assemblies
Connector Type: (The connector pin-outs differ) ....
IEEE488/ANSI MC1.1; 24 pins, accepts a ribbon cable. ....
IEC 625-1, MIL-C-24308; 25 pins
Connectors are tested to MIL-STD-202E (Environmental).

IEEE488 has a Maximum cable length of 20 meters, or 2 meters per device - which ever is less.
IEEE488 devices may be connected in either a Star or Linear fashion [Device-to-Device].

GPIB Cable Manufacturers
Belden
Keithley {Single/Double shielded GPIB Cables, IEEE-488 to Centronics cable adapter}
INES Test and Measurement
Normal cable lengths; 0.5 meters, 1m, 2m, and 4 meters.

Metric threads are Black, English threads are Silver; the two will not mate together.